“Addition of Navy Reports to Calibration Report Utility”

# TABLE OF CONTENTS

I.	TABLE OF CONTENTS	2

II.	REVISION LOG	3

III.	PROCEDURE STATUS	3

IV.	REFERENCES	3

V.	PURPOSE	4

VI.	SCOPE	4

VII.	PERSONNEL TRAINING AND QUALIFICATION	4

VIII.	APPARATUS AND EQUIPMENT	4

IX.	CURRENT PROCEDURE	4

X.	PROPOSED CHANGES	9

A.	Modification of DB Configuration Files.	9

B.	Navy Report Drop Down Option.	11

C.	Unit Type Selection Prompt.	14

D.	Navy Report Front Panel	17

E.	Report Value Population	19

F.	List of Checked Out Files	26

XI.	NEW PROCEDURE	27

# REVISION LOG

| Revision: 1.0 |  | Effective Date: |  |  |
| --- | --- | --- | --- | --- |
| Section | Justification |  |  | Changes |
| Impact of above change(s) on other documents: |  |  |  |  |

# PROCEDURE STATUS

Current Status is Draft.

# REFERENCES

https://mysite.jci.com/personal/ckingco_jci_com/Documents/Forms/All.aspx?e=5%3A77f7282bb6904346b51e62ff5a24c19f&sharingv2=true&fromShare=true&at=9&CID=1c1881ae%2D624f%2D4c80%2D938f%2D8138e2b85e1d&RootFolder=%2Fpersonal%2Fckingco%5Fjci%5Fcom%2FDocuments%2FDesktop%2FNavy%20Reports&FolderCTID=0x0120008D3279C5D6CF504DBE9B907E73704B91

https://leplb0040.upoint.ap.alight.com/ah-angular-afirst-web/ - /web/jci/homepage

# PURPOSE

To document the current process for creating a Navy Report and show the proposed changes to integrate the process into the current Calibration Report Utility.

# SCOPE

Scow current and future proposed Navy Report Process.

# PERSONNEL TRAINING AND QUALIFICATION

- Excel Experience

- Knowledge of Calibration Report Utility

- Access to OneDrive

- NI LabVIEW

# APPARATUS AND EQUIPMENT

- Microsoft Excel

- LabVIEW 2015 Runtime

- Calibration Report Utility.exe

- Microsoft One Drive

- NI LabVIEW 2015 SP1

# CURRENT PROCEDURE

- An email from someone, requesting a calibration report for instruments will come in.

  - Example:

Subject: FW: Labview TR Info - 570 Production TR24-0043

Can I please get the instrumentation calibration report for the instruments on the attached map as well as the 1200 SR instruments. Please use TR 24-0043 and run 15.01 taken on 7/8/24.

  - Important information is highlighted.  This information will be used to get the proper information as well as build the header.

    - Size of 570 Ton

    - Test Request TR24-0043

    - System 1200 SR

    - Date of 07/08/2024

- Run Calibration Report Utility.exe 100-21-039 1.3.0.12

- Log in and load the system using the Unit specified in the email.  In this case, the unit specified is 1200 SR or WC1200.  Also select the appropriate date.  In this case the date of 07/08/2024 will be selected.

- Next:

  - Select “Combined Report”

  - Check “Include Inactive HW Tags” & “Include Engineering Tags”

  - Press “Generate Cal Report”

- This will generate a CSV report and open up the location in Windows explorer.

- The report then can be opened in Microsoft Excel.

- From this point there are templates that are used to create the report.  Those templates can be found here: https://mysite.jci.com/personal/ckingco_jci_com/Documents/Forms/All.aspx?e=5%3A77f7282bb6904346b51e62ff5a24c19f&sharingv2=true&fromShare=true&at=9&CID=1c1881ae%2D624f%2D4c80%2D938f%2D8138e2b85e1d&RootFolder=%2Fpersonal%2Fckingco%5Fjci%5Fcom%2FDocuments%2FDesktop%2FNavy%20Reports&FolderCTID=0x0120008D3279C5D6CF504DBE9B907E73704B91

- In this instance, the template will be 570 Ton.  Open the “570 Ton” folder.

- Create a new folder in this location and give it the date for the name.  Place the generated CSV Report in this location.

- In the “570 Ton” folder, is the template for the report, “570 Ton Production_Template_RevD”  Place this file in the new folder created in the previous step.  Rename it to “570 Ton Production_Date”

- The next step will be to open the modified report and add in data.

  - Open, copy, and paste in the instrument map into the instrument map tab of the report.

  - Open, copy, and paste in the generated combined report into CalReport tab of the report.

- Next steps are to populate the “Report Worksheet”

  - Select the “Report Worksheet” Tab.

  - The 1st section to fill out is the header details.

    - “Unit Type” 570 Ton

    - “Test Request Number” TR24-0043

    - “Instrument Map Date” ‘7/9/2024

    - “JADEC Test Facility” WC1200 SR

    - “Calibration Request Date” 7/9/2024

  - Next the Serial Numbers for “DAQ2 RTD Temperatures” need to be found and populated.

    - Use the “Instrument Map” to match up the “LabVIEW Tag Names” to the “ENG Tag Names”

    - From here, the “DAQ Driver” and “DAQ Channel” can be found for the Tag.

    - “DAQ Driver” and “DAQ Channel” and the “CalReport” Tab Data, match up and find the “Instrument Serial Number” and list it in the proper column.

  - Repeat these steps for:

    - “DAQ3 RTD Temperatures”

    - “Omega Pressure Transducer”

    - “Rosemount Delta P Transducer”

    - “Rosemount Magmeter”

  - For the “Xitron Power Analyzer” portion, the information can be found towards the top of the “CalReport”

    - Copy in the “Serial Number” and the “System Name”

  - Lastly, fil ouy “Prepared by” and “Date”

- Verify the checks are all green on the right side of the “Report WorkSheet”

- Switch to the “Report Template” Tab

- The report should be fully filled out.

  - Full Report

# PROPOSED CHANGES

## Modification of DB Configuration Files.

- Check out of configuration files of various locations for modifications.

  - C:\_Projects\_LV2015\TrendSafe\SDC\Configuration Data\ConfigurationDB.UDL

    - Modify Option 1 Server Name to be j030m21d

    - Modify Option 2 to “Use Windows NT Integrated security”

    - Modify Option 3 to “SDC_Configuration”

    - Connection is Tested for success.

  - C:\_Projects\_LV2015\TrendSafe\SDC\Configuration Data\StroageDB.UDL

    - Modify Option 1 Server Name to be j030m21d

    - Modify Option 2 to “Use Windows NT Integrated security”

    - Modify Option 3 to “SDC_Storage_Federal”

    - Connection is Tested for success.

## Navy Report Drop Down Option.

- Check out of C:\_Projects\_LV2015\TrendSafe\SDC\GUI\Classes\CalibrationReport_class\Stand Alone Calibration Report-UI.VI for Drop Down Modifications.

- Current Stand Alone Calibration Report-UI.VI look.

- Current “Report Type” Options.

  - Combined Report

  - Module Report

  - Sensor Report

  - Power Analyzer Report

- Proposed Addition of “Navy Report”

  - Check out of C:\_Projects\_LV2015\TrendSafe\SDC\GUI\Classes\CalibrationReport_class\CalReportType—Enum.ctl for addition of “Navy Report”

    - Current Type Def Report Types

    - Combined Report

    - Module Report

    - Sensor Report

    - Power Analyzer Report

    - Modified Type Def Types

  - Modification of Stand Alone Calibration Report-UI.VI Case Structure to add new “Report Type”

    - Current Case Structure

    - Temporary Modified Case Structure

  - Stand Alone Calibration Report-UI.VI after Type Def Update and addition of “Navy Report”

## Unit Type Selection Prompt.

- Source for Unit Type selection for Unit Type Selection Prompt.  Unit Type will be derived from filenames of configuration files in a folder location.  This way when a report is added or removed, no code change is needed.

  - Creation of “Navy Report” Folder in location C:\_Projects\_LV2015\TrendSafe\SDC\Configuration Data\JADEC_CAL_REPORT_UTILITY

  - Creation of “Unit Types” Folder in Location C:\_Projects\_LV2015\TrendSafe\SDC\Configuration Data\JADEC_CAL_REPORT_UTILITY\Navy Reports

  - Creation of “Unit Type” Configuration Files.  File will be .ini

    - 110 Ton.ini

    - 200 Ton Tiper.ini

    - 375 Ton HESC.ini

    - 450 Ton.ini

    - 500 Ton.ini

    - 520 Ton.ini

    - 570 Ton.ini

    - 1100 Ton.ini

  - File Tytpe will need to:

    - Hold Template information.

    - Be able to easily add or remove data.

    - Include the sections of the report.

    - Include the data to search for.

    - Include the data to insert.

  - File Format

[Report Section 1]

“ENG Tag to Search For”=”Instrument Title

“ENG Tag to Search For”=”Instrument Title

“ENG Tag to Search For”=”Instrument Title

[Report Section 2]

“ENG Tag to Search For”=”Instrument Title

“ENG Tag to Search For”=”Instrument Title

“ENG Tag to Search For”=”Instrument Title

…”and So forth”

- VI Creation for Unit Type Selection Prompt.

  - Requirements

    - Layout

      - Drop down for Unt Type

      - Okay Button

      - Cancel/Exit Button

      - VI Run Time Position of Centered on the Primary Monitor.

    - Size

      - Smaller than Stand Alone Calibration Report-UI.VI Front Panel.

    - Color Scheme

      - Match Stand Alone Calibration Report-UI.VI

    - Operation

      - Prompt will launch when “Generate Cal Report” is pressed on Stand Alone Calibration Report-UI.VI and Navy Report is selected as the “Report Type”

      - User will be able to select the unit type and press “Continue” to proceed or “Exit” close out of the prompt and not proceed with the generation of the cal report.

      - Only unit types from the folder are displayed and allowed.

- Location of Selection Prompt.

  - C:\_Projects\_LV2015\TrendSafe\SDC\GUI\Classes\CalibrationReport_class

  - Create a new VI by right clicking o CalibrationReport.lvclass New  VI

- Using the requirements, an appropriate front panel is created.

- Initial draft of Selection Prompt for Unit Type.

- Unit Type Selection Prompt with possible Unit Types.

- An intial draft of a block diagram wil need to be created to allow test functionality.  This functionality will:

  - Pull in the config files to populate the drop down.

  - Use an event structure for output of the user selection or whether Exit is pressed.

- Check Out of C:\_Projects\_LV2015\TrendSafe\SDC \JADEC-PC.lvproj to allow for addition of  Unit Type Selection Prompt to be added to the project and CalibrationReport.lvclass library.

## Navy Report Front Panel

- Open Create Sensor Calibration Report.vi and save a copy.  Call the copy Create Navy Report.vi

- Modify the front panel to resemble the Navy Report Document.

  - Header

    - Title – Johnson Controls Calibration Report.

    - Unit Type – Derived from Unit Type Selection.

    - Test Request Number – Derived from the Loaded System.

    - Instrument Map Date – Derived from the Loaded System Date.

    - JADEC Test Facility – Derived from the loaded System.

  - Report Body

    - Serial Number – Unit Serial Number.

    - Instrument Title – Pulled from the .ini for the Unit Type and inserted upon finding the ENG Tag Name, also supplied by the .ini for the Unit Type.

    - HW Tag Name – Unit HW Tag Name.

    - Calibration Dates – Dates of Calibration.

      - Next – Calibration Due Date.

      - Last – When the Calibration was last done.

      - Status – Compares the date and the next calibration to determine if the unit is within calibration.

- Final Design Layout

- Check out C:\_Projects\_LV2015\TrendSafe\SDC\GUI\Classes\CalibrationReport_class\CalibrationReport.lvclass to allow addition of Create Navy Report.VI to the class library.

## Report Value Population

- Stand Alone Calibration Report-UI.VI Modifications

  - Add a new Case to the Case Structure in Stand Alone Calibration Report-UI.VI, under "", "smUserAction", and name the case Navy Report

  - Add the Unit Type Selection Prompt inside the case, in line with error and use the unit type output to route to another case structure.

  - Inside this case structure, create a case, “Exit” to be empty, in the event that Exit is pressed on the unit type selection prompt.

  - The other Case will be the default case.  In the default case, route in the GUI Loop Cluster and unbundle out:

    - Active TR

    - Reference Tag List

    - System Name

    - Time Stamp

  - Route the unbundled values from the previous step to the newly created Create Navy Report.vi.

  - The case should look like the following:

- Create Navy Report.vi Modifications

  - FullBlock Diagram modification is needed to obtain the desired report values.

  - Message Cases

    - Clear Report – sends a blank array to the table to clear the table.

    - Populate – Populates Report Title, Date, Instrument Map Date, and Report.

    - Get Systems – Get Systems, obtains the systems and populates the shift register.

    - Obtain All Active Tags – Obtains the active tags from all systems.

    - Obtain Power Analyzer Info New – Obtains the info about the power analyzers.

    - Get Config File Data – Opens the Unit Type Config File and obtains the information.

    - Build Report – Searches the active tags and power analyzer info, using the config file data, to obtain the data needed.  Then the data is put into the report.

    - Format Report – Sizes the table to the report size and moves the created by and date to the bottom of the table.

    - Print Report – Priints the front panel.

    - Exit – closes the message handling loop.

    - Error, Default – Handles Errors.

## List of Checked Out Files

- C:\_Projects\_LV2015\TrendSafe\SDC\Configuration Data\ConfigurationDB.UDL

- C:\_Projects\_LV2015\TrendSafe\SDC\Configuration Data\StroageDB.UDL

- C:\_Projects\_LV2015\TrendSafe\SDC\GUI\Classes\CalibrationReport_class\Stand Alone Calibration Report-UI.VI

- C:\_Projects\_LV2015\TrendSafe\SDC\GUI\Classes\CalibrationReport_class\CalReportType—Enum.ctl

- C:\_Projects\_LV2015\TrendSafe\SDC \JADEC-PC.lvproj

- C:\_Projects\_LV2015\TrendSafe\SDC\GUI\Classes\CalibrationReport_class\CalibrationReport.lvclass

# NEW PROCEDURE

- Run “Calibration Report Launcher.vi”

- This will open the calibration report launcher.

- Use the Icon on the upper left to Log In.

- Check “Include Inactive HW Tags” and “Include Engineering Tags”

- Click “Load System”

- This will open the “System Selection”

- Select the desired system, Date, and the press okay.  For this example, we will use WC1200 and the current Date and Time.

- The system will then load the selected system.

- Select “Navy Report” From the Report Type drop down.

- Click “Generate Cal Report”

- The Unit Type Selection will display.  Select the Unit Type from the drop down and then select Continue.

- This will then bring up the Navy Calibration Report Window with the Report.  From here, the “Print” button can be used to print the report of the “x” can be used to close the report, without printing.

- Print Option will bring up a “Save Print Output As” Prompt. From here, a location can be selected to save the PDF Report File.

- This report file can then be opened with a PDF viewer.